Fixed-clamp-on ultrasonic flow meter / LRF-2000SW


· LRF-2000S is one of the best-selling flow meters. It enjoys high popularity because of its economical price and reliable function. Installation is quite easy and operation is convenient.

Product Selection Manual

1.Features
· High Accuracy measuring 
Linearity: better than 0.5%, Repeatability: better than 0.2%, Accuracy: better than ±1% 
Wide Measuring Range 
Several types transducer for selection, measuring pipe size from DN15mm to DN6000mm
· Chinese or English menu 
Different version UFM can support Chinese or English display, easy to operate
· High Reliability 
Adopt low voltage, multi-pulse technology to improve accuracy, useful life and reliability
· Strong Anti-interference 
Dual-balance signal differential receiver/driver circuit to avoid interference of converter, TV tower, high voltage line etc.
· Powerful Recording Function 
Automatically record the following data: 1. The totalizer data of the last 512 days/128 months/5 years 2. The time and corresponding flow meter of the last 64 times of power on and off events 3. The working status of the last 32 days
2.Main Feature
	 Item
	 Performance and Parameter

	 Convertor
	 Principle
	 transit-time ultrasonic flowmeter

	
	 Accuracy
	 Flow: over 1%-2%(battery-powered supply over ±2% )，Heat: over ±2%，Repeatability: over 0.2%

	
	 Display
	 2×20 character LCD with backlight, support the language of Chinese, English and Italy Convertor dimension:170*180*56cm

	
	 Operation
	 Support 16 key light touch keyboard and the Simulation keyboard software

	
	 Signal Output
	 1 way 4~ 20mA output, electric resistance 0~ 1K，accuracy0. 1%

	
	
	 1 way OCT pulse output( Pulse width 6 ~ 1 0 0 0 ms，default is 2 0 0 ms)

	
	
	 1 way Relay output

	
	 Signal Input
	 3 way 4~ 20mA input, accuracy 0. 1%，acquisition signal such as temperature, press and liquid level

	
	
	 Connect the temperature transducer Pt100, can finish the heat/energy measurement

	
	 Data interface
	 Insulate RS485 serial interface, upgrade the flow meter software by computer, support the MODBUS, METER-BUS etc.

	
	 Other Function
	 Automatically remember the previous flow/energy data of 512 days,  128 months and 10 years

	
	
	 Automatically remember the previous flow/energy data of 512 days working condition and check it

	
	
	 Automatically remember the previous 32 times power on/off time and flow rate

	
	
	

	 Special Cable
	 Twisted-pair cable, generally, the length under 50 meters; Select the RS485 communication. The transmission distance can over 1000m.

	 Pipe Condition
	 Pipe Material
	 Steel, Stainless steel, Cast iron, Copper, Cement pipe, PVC, Aluminum, Glass steel product, liner is allowed.

	
	 Pipe Diameter
	 25---6000mm

	
	 Straight length pipe
	 Transducer installation should be satisfied: upstream10D, downstream 5D, 30D from the pump.

	 Measuring Medium
	 Type of Liquid
	 Water (hot water, chilled water, city water, sea water, waste water, etc.);  Sewage with small particle content; Oil (crude oil, lubricating oil, diesel oil, fuel oil, etc.);  Chemicals (alcohol, acids, etc.); Plant effluent; Beverage, liquid food; Ultra-pure liquids; Solvents and other liquids

	
	 Temperature
	 Standard Transducer: -30--90℃; High-temperature transducer: -40--160℃

	
	 Turbidity
	 No more than 10000ppm and less bubble

	
	 Flow rate
	 0--±12 m/s

	 Working Environment
	 Temperature
	 Main Unit: -20--60℃

	
	
	 Flow transducer: -40--160℃; 
Temperature transducer: according to the customer’s demands.

	
	Humidity
	Main Unit: 85% RH; Flow Transducer: Ip67

	Power Supply
	DC  24V or AC  85-264V

	Power Consumption
	1.5W



3.Optional Main unit 
    According to different installation condition,there are three models separated ultrasonic flowmeter for selection:
	                  Wall-mount
	 Panel-mount
	Explosion-proof

	 [image: image1.jpg][

—_—

" [cese] =
[sc5el
s

q |axo
i
4 = .




	 [image: image2.jpg]—




	 [image: image3.jpg]




	Wall mounted LRF-2000SW
Indoor or instrument box installation
Dimension:170*180*56mm
Power supply:DC8~36V or AC85~264V
	Panel mounted LRF-2000SS
Instrument Cabinet installation
Dimension:170*180*56mm
Hole-opening dimension:156*76mm
Power supply:DC8~36V or AC85~264V
	Explosin-proof LRF-2000SD
Explosin-proof  installation
Dimension :298*298*110mm
Power supply:DC 8~36V or 
AC85~264V



4. Optional transducer 
	Optional Transducers
	Model
	Measuring Range
	Liquid Temperature
	Accuracy

	Clamp-on
	TS-2(small size)
	LRF-2000SNB1
	DN15-100mm
	-40C°∼90C°
	Better than

±1%

	[image: image4.jpg]



	TM-1(medium size)
	LRF-2000SNB2
	DN50-1000mm
	
	

	
	TL-1(large size)
	LRF-2000SNB3
	DN300-6000mm
	
	

	High-tem clamp-on
	TS-2-HT(small size)
	LRF-2000SNB4
	DN15-100mm
	-40C°∼160C°
	

	[image: image5.jpg]



	TM-1-HT(medium size)
	LRF-2000SNB5
	DN50-1000mm
	
	

	Insertion
	TC-1(standard)
	LRF-2000SNC1
	DN80-6000mm
	-40C°∼90C°
	

	[image: image6.jpg]



	TLC-2(lengthen)
	LRF-2000SNC2
	
	
	

	Prover Section
	Prover Section
	LRF-2000SNG
	DN15-1000mm 
(see appendix 2)
	-40C°∼90C°
	Better than

±0.5%

	[image: image7.jpg]



	
	
	
	
	


